Vision pipeline issue:

Video processing box to console:
- Problem with the left display, very likely caused by the SDI cable
- If SDI cable of the right display is disconnected, both displays go down
**Why the hell is the camera input coupled on right input cable!?
- switching left and right SDI input doesn't help
- unpluggin/switching SDI cables on processing box doesn't affect the video on console
2D display = only right camera footage displays on both screens
3D display = each camera footage works individually

Console to computer:
- nothing shows up on the display (both left and right)
- the converter loses power after the display option shows up (it's like the power cable just gives up halfway)
** once the converter's power cable is connected, the display option shows up in the setting whether the SDI cable is connected to the console or not
- HDMI to SDI converter works, but nothing on displays

video processing box to computer
- left SDI output from video processing box gives SMPTE
- right SDI output from video processing box works fine

No footage on left display – resolved
reconnect the output from camera control to video processing box
- new issue: lost the endoscope footage on Trinitron
No footage on Trinitron – resolved
reconnecting the output from console to Trinitron

Current issue:
- no connection from PC to console
- the image on the Trinitron screen is off the edge, images on console displays are normal

MTM console video input
Resolution: 720x576 @ 50.00 Hz
Analog: NTSC
Issue:
Luma signal: screen horizontally out of sync + the whole screen moves out of place occasionally
Chroma signal: RGB color is completely messed up
If the video processing box is connected and turned on, MTM won't accept the connection from PC

custom refresh rate and resolution doesn't work. Horizontal sync still out of sync
custom resolution might work, although not available by default: 640x480 @ 59.94 Hz

for S-video cable from video processing box:
doesn't work on left display channel (green screen). Works on right display channel with nightvision-like visual.

Suspects : S-video to coax could be the cause
One of the possible solutions:
design a filter that takes PC output as input and fix the image distortion from there.

Oscilloscope readings:
voltage from luma signal from pc is higher than from sync box.
Voltage from chroma signal from pc is lower than from sync box.
Issue fixed: get good S-Video cable. :)
note: to get both displays to work, both S-Video inputs must be connected. (still have some issue with left display)

next task: stereo camera simulation with overlay